Lensmaker/SL² - Limit Switch Adjustment Procedure.        

Date Composed: 17/11/95   
Date Modified:
25/07/96

Prior to making adjustment to the limit switch, vacuum all foreign material, debris, etc. from the limit switch and it's actuator

.  

To properly adjust the limit switches on the slides, follow these instructions:

Y - axis  

The Y - negative limit switch (towards the operator) must actuate at 1/4 turn of the Y - axis leadscrew from the front (towards operator) mechanical hardstop.

The Y - positive limit switch (away from operator) must actuate at 1/2 turn of the Y - axis leadscrew from the rear (away from operator) mechanical hardstop.

X - axis 

The X - positive limit switch (away from the chuck) must actuate at 1/4 turn of the X - axis leadscrew from the rear (away from the chuck) mechanical hardstop.

The X - negative limit switch (toward the chuck) must actuate at 1/4 turn of the X - axis leadscrew from the front (toward the chuck) mechanical hardstop.

       NOTE: Once these switches are adjusted, ensure that they are securely fastened and  

                        unable to pivot on their mounts.
